# SOP pre-processing Methylight raw data qWID-GYN

Last updated: 14.12.2022, Charlotte Vavourakis

Save the .eds file in the ZIDshare

Make a copy of the .eds file to your project folder and experiment subfolder in:

\Dropbox\eutops\data\raw_data\methylight\

Generate a samplesheet using the Shiny

- Make sure R is installed and in your environment path. Libraries needed: shiny, readxl, stringr and dplyr.

- Navigate to \Dropbox\eutops\scripts\methylight\ML_WIDqGyn\shiny

- For Windows: Double-click GeneratePlatelayoutGYN.bat (you could make a shortcut to this file e.g. on your Desktop)

- For Mac/Linux: Double-click GeneratePlatelayoutGYN.sh

- Follow instructions in the browser.

- Please note that .bat files are specific for Windows OS.

Save the resulting sample sheet, alongside the raw data (.eds file) both on the ZIDshare and the Dropbox

- !!!!!!!!!! Make sure to do this BEFORE editing an .eds file !!!!!!!!!!

Using the Quantstudio Design & Analysis software 2.6.0, extract, normalize and export run results

- Open .eds file

- Import your sample sheet under Plate setup > … > Import Plate Setup

- Press ok, when prompted to overwrite Plate Setup

- Under Plate setup, Add the correct Quantities for the standards in the Sample list (example given for gBlocks)

- Press Analyze

- Under Quality check inspect the Amplification Plots

- You can select specific wells, checking minimally Standards, Positive and Negative controls.

- Delta Rn should be normalized on the Y-axis using the passive Reference dye ROX, showing a scale from 0 to 1

- Under Actions, Primary Analysis Settings, Set the Thresholds, and press Save:

  - COL2A1 = 0,12

  - ZSCAN12 (FAM) = 0.17

  - All other targets = 0.04

- Under Actions export the result files (default setting)

- Save them to your project folder and experiment subfolder in:

\Dropbox\eutops\data\raw_data\methylight\

- Close and save the .eds file

Calculate PMR, WID-qEC and WID-qCIN using Shiny

- Navigate to \Dropbox\eutops\scripts\methylight\ML_WIDqGyn\shiny

- For Windows: double-click CalculatePMRGyn.bat (you could make a shortcut to this file e.g. on your Desktop)

- For Mac/Linux: double-click CalculatePMRGyn.sh (you could make a shortcut to this file e.g. on your Desktop)

- Follow instructions in the browser.

Please update “experiment_log” \Dropbox\eutops\data! For larger cohorts and studies